## Slide 1
WakeMed – CSIC Data Quality Overview
Data volume
Work orders by year
assets
99k
CM
69k
work orders
200k
62k
PM
40k
Other
WO coverage
Jan 2018 May 2025
Chart: 
| 2018 | 2019 | 2020 | 2021 | 2022 | 2023 | 2024 | 2025
Series 1 | 14500 | 15232 | 25030 | 29254 | 28896 | 30182 | 32340 | 12039

## Slide 2
WakeMed – Usability
Top 5 modalities (across all CSIC participants)
Patient Monitoring Systems
Patient Care and Exam
Infusion Pumps
Respiratory
Surgery
Top 5 modalities across all CSIC participants. May differ from top 5 modalities for Wakemed
Low usability for MTBF because all assets require at least 1 CM event for MTBF calculation, and most assets require at least 2 CM events
% WOs with matching asset = 1 - % orphaned work orders. Orphaned work orders = work orders whose asset ID is not found in the asset dataset
Orphaned work orders don’t have a matching asset, so the asset’s modality is unknown. % orphaned work orders and % of WOs with matching assets cannot be calculated for a subset of modalities / makes / models
 | Top 5 modalities |  | All modalities | 
 | Wakemed1 | All CSIC participants | Wakemed | All CSIC participants
% assets with age | 100% | 96% | 100% | 97%
% assets with PMs | 10% | 36% | 15% | 35%
% assets with CMs | 24% | 15% | 27% | 17%
% assets with MTBF2 | 16% | 7% | 16% | 8%
% WOs with matching asset3 | n/a4 | n/a4 | 80% | 77%
% WOs with cycle time | 84% | 41% | 83% | 60%

## Slide 3
WakeMed – Work Order Types
 | # WOs | % WOs
Preventive Maintenance - CE | 29,431 | 14.7%
Planned Maintenance - CE | 19,451 | 9.7%
Incoming Inspection | 12,682 | 6.3%
PM unscheduled | 140 | 0.1%
New Equipment Inspection | 12 | 0.0%
Total PM | 61,716 | 30.8%
 | # WOs | % WOs
Incoming Inspection | 12,682 | 6.3%
Asset Onboarding - CE | 6,218 | 3.1%
Administrative Time | 5,268 | 2.6%
Safety Test Non Hosp | 4,328 | 2.2%
Asset Disposal - CE | 2,085 | 1.0%
Project | 1,896 | 0.9%
Recall - CE | 1,114 | 0.6%
Biomed Loaner | 688 | 0.3%
Remove from Service | 685 | 0.3%
Rad Physicist Survey | 681 | 0.3%
Scopes and Instruments | 613 | 0.3%
NPO Short Term Demo/Rental | 612 | 0.3%
Other | 589 | 0.3%
Middle Room - Raleigh | 536 | 0.3%
Alert/Recall Inspection - CE | 506 | 0.3%
Checkout Disp (Accessory) | 471 | 0.2%
Physics Inspection | 294 | 0.1%
Other1 | 587 | 0.3%
Total Other | 40,046 | 20.0%
1. Gemba - Service, Education / Inservice, Incident Investigation, Software/Firmware Upgrade, Nuvolo Project, Raleigh OffSites, Vulnerability Remediation, Dept. Rounds, Device Verification, Cary Middle Room, New Equipment Inspection, Tech Support, (Blanks), Upper Southern Route, Lower Southern Route, Equipment Distribution, Calibration Testing, Fireman
Preventive maintenance
Other
 | # WOs | % WOs
Service-Checkout-Repair | 53,639 | 26.8%
Corrective Maintenance - CE | 44,768 | 22.4%
Incident Investigation | 104 | 0.1%
Tech Support | 4 | 0.0%
Device Related Incident | 1 | 0.0%
Total CM | 98,516 | 49.2%
Corrective maintenance

## Slide 4
CSIC Definitions
KPI | Definition | Observations
# of assets | # of asset IDs for a given modality, make, customer, etc. | Based on de-duplicated counts (i.e. unique asset IDs)
# of work orders | # of work order IDs for a given asset, modality, make, etc. | Based on de-duplicated counts (i.e. unique work order IDs)
Work order timeframe | Time between first and last work order | Based on reported date for each work order
Asset age | Time since asset install date | Asset age = NULL for assets with unavailable or invalid install date Install dates before 1970 or after today’s date are deemed invalid
% work orders with matching asset | % of all work orders for assets found in the asset dataset | % WOs with matching asset = 1 - % orphaned work orders Orphaned work orders = work orders whose asset ID is not found in the asset dataset
MTBF (mean time between failure) | Time between failures for a given asset | “Failure” is defined as any corrective maintenance event Time between CM events (“failures”) is based on each CM’s reported date Only includes time between known CMs. Work orders opened prior to asset install date or after asset retirement are not included For aggregations (e.g., MTBF of a certain make / model / modality), MTBF is calculated as the sum of time between failure events divided by the sum of failure events
WO cycle time | Time between work order open and closed date for a given work order | Only work orders with status “Closed Complete” are included in the calculation